## Slide 1: Installation and Commissioning Plans for the jFEX
(Katharina Bierwagen)
Uli/Julio

## Slide 2: Agenda
Introduction
jFEX Status
Hardware Availability
Surface Test Facility
Installation
Commissioning
Summary

## Slide 3: Introduction
jFEX identifies jet candidates and calculates and for each BC
|η|< 4.9 with different granularities
6 modules to be built
Hardware details
  ATCA board
  4 Xilinx UltrascalePlus FPGAs per module
    XCVU9P-2FLGA2577E
    Up to 120 MGTs per FPGAs
  24 MiniPOD: 20 RX + 4 TX
  Modular design
  Mezzanines
    Power
    Control
jFEX functionality – Block Diagram

## Slide 4: jFEX Status
jFEX Final Prototype
Final Design Review took place end of May
  Firmware: May 29
    https://indico.cern.ch/event/728398/
  Hardware: May 30
    https://indico.cern.ch/event/728400/
Green light for producing the jFEX Pre-Production modules on June 12
Ongoing PCB production
Order of all components placed
Expect delivery from assembly company: beginning of October
Firmware
  Ongoing integration of infrastructure firmware (MGTs, board control, IOs and etc.) and algorithms

## Slide 5: Hardware Availability (Plan)
jFEX First Prototype – Currently at CERN (STF)
  Equipped with one UltraScale FPGA
jFEX Final Prototype – Currently in Mainz
  Equipped with four UltraScale+ FPGAs
jFEX Pre-Production Module – Available from November 2018 onwards
  Pre-Production module will be used as spare
jFEX Production Modules – Available from May 2019 onwards
  9 (+ 1) modules
    6 to be installed
    3 (+ 1) spares

## Slide 6: Surface Test Facility (1)
No hardware changes are foreseen from prototype to production module
Initial integration tests with other systems can start at anytime with Final Prototype
  And with Pre-Production by November
Combined tests with TREX schedule for beginning of September (in Heidelberg)
Combined tests with ROD (Hub) to take place soon
Final Prototype = Pre-Production Module = Production Module

## Slide 7: Surface Test Facility (2)
Setup ready for tests with jFEX First Prototype and FTM since July
  Tests of optical links, TTC clock reception, IPBus and etc. were performed
  jFEX Final Prototype is currently in Mainz, but will be shipped back to CERN soon
What can and cannot be done elsewhere?
  Will we be able to perform tests with some calorimeter boards?
  How/what will be the DCS setup on the STF?
jFEX
FTM
“TTC system”
jFEX setup on the STF

## Slide 8: Surface Test Facility Pre-Production Module
Validation of Pre-Production module will take place at the STF
    Smoke tests and basic validation of board optical/electrical links (IBERT) to be done in Mainz
  Validation of all optical/electrical inputs/outputs simultaneously (IBERT and formatted data)
  Power consumption
    Ripple measurement
    Thermal dissipation
  LHC clock reception
    MGTREF clock jitter measurement
  Validation of the backplane high-speed links
    Readout and TTC data
    Latency measurement
    IPMC
Firmware consolidation
  Tests with integrated firmware (infrastructure and algorithms)
  Latency measurement
  Combined tests with other parts of the system

## Slide 9: Surface Test Facility Production Module
Initial production modules acceptance tests will take place in Mainz
  Smoke tests and basic validation of board optical/electrical links (IBERT)
Finalization of production modules acceptance tests on ATCA crate with HUB/ROD and/or FTM
  Same test list of the Pre-Production module
Full slice test from Latome/TREX via jFEX to Topo
  As much as possible inputs
    Formatted data
  With all possible infrastructure parts
    Hub/ROD, FELIX, IPBus, DCS and etc.
  Final evaluation on power consumption and cooling with closed shelf
Integration with “L1Calo STF partition”

## Slide 10: USA15 Installation (1)
jFEX connections
  LATOME
  TREX
  FOX
  Hub/ROD
    FELIX
  Topo
    Installation schedule for December 2019
Initial tests
  Verification of board functioning after transport
  Establishment of board communication and control
    IPbus and DCS
  Integration tests with HUB/ROD
  Combined tests with calorimeters
Ideally, similar to what was done in STF during 2018 and 2019
To be installed before jFEX
To be defined?
1 fibre via front panel
4 fibres through the backplane
TTC data, LHC clk, IPbus and etc.
ATCA
1 fibre via front panel
RTM
HUB/
ROD
PC*
Topo
TileFOX E or TileFOX F
jFEX
FELIX
* To be used during standalone tests at the beginning of installation

## Slide 11: USA15 Installation (2)
Starts in November 2019
  Followed by 3 months re-integration & commissioning
  What is the installation sequence?
    Will the crate be already in USA15?
      With all FOX connections done?
    Or will it be transported from the STF with all boards?
    When will the Hub/ROD be installed?
  Will we take care of transportation and fibre connection ourselves?
  Should we perform tests with the Prototype/Pre-Production module in USA15 before installation?

## Slide 12: Commissioning
Validation of (FOX) input mapping
Integration with ATLAS partition
Online monitoring
Combined tests with calorimeter
Timing calibrations
Pre-beam commissioning with cosmics
Commissioning with beam
Validation with bit-wise simulation
Cross-comparison with legacy system

## Slide 13: Summary
Ongoing production of the jFEX Pre-Production module
Pre-Commissioning to take place during 2018 and 2019 at the STF
  Setup ready on the STF for jFEX
  Finalization of production modules acceptance tests
Installation in USA15 scheduled for November 2019

## Slide 14: jFEX Schedule